WIRO 8-AXIS CONTROLLER

USER'S GUIDE

Please submit corrections and suggestions to the Observatory Engineer.

| Revision | Date By Description |
| --- | --- |
| 0 | 27Sep04 JSW 35 pages plus 51 page appendix |
| 1 | 30Jan07 JSW Page 33: alter comments regarding microstep switch settings and the "shutdown" command. Page 34: switch LED's not yet implemented and 6K8 Ethernet LED must blink once. |

WIRO 8-AXIS CONTROLLER USER'S GUIDE

Table of Contents

1.  Purpose of the 8-Axis Controller...........................4

A.  Why Have an 8-Axis Controller?...........................4

B.  What Can this 8-Axis Controller Do?......................4

2.  Scope of this Manual.......................................5

3.  Other Reference Material...................................6

4.  Basic Description..........................................6

A.  Layout and Cabling.......................................6

1.  WIRO 8-Axis Controller Box Layout......................9

2.  WIRO 8-Axis Controller Cassegrain Focus Cabling.......10

3.  WIRO 8-Axis Controller Prime Focus Cabling............11

B.  Functions...............................................12

1.  Compumotor 6K8 "Limits/Home" Schematic................13

2.  Compumotor 6K8 "Triggers/Outputs" Schematic...........14

3.  Compumotor 6K8 Sample Input Connections...............15

4.  Compumotor 6K8 Sample Output Connections..............16

5.  Compumotor 6K Series Commands.........................17

6.  IMS Model IM483 Microstepping Driver Quick

Reference Card........................................18

C.  Electrical Configuration................................20

1.  Typical Motor Drive and Power Supply Configuration....21

2.  Telescope Backplate Connector.........................22

3.  Dome Optics Bench Connector...........................23

4.  Telescope Top-End Connector...........................24

5.  Channel Selector Switch for Motors 4 & 5..............25

5.  Interfacing the 8-Axis Controller to an Instrument........26

A.  Hardware Issues.........................................26

1.  Interface Specifications..............................26

a.  Stepper Motor Selection.............................26

b.  Sensor Selection for "Home", "Limit

or "Trigger" Inputs.................................26

c.  "Outputs" Capability................................26

d.  Power Capability....................................26

2.  Applications..........................................26

a.  Mounting Hardware on the Instrument.................26

b.  Cabling the Controller to the Instrument............28

1.  Wire Gauge and Length Restrictions................28

2.  Twisted Pair Wiring...............................29

3.  Shielding.........................................29

4.  Wiring Diagram for Circular Connectors............31

5.  Wiring Diagram for Elco Connectors................32

B.  Software Issues.........................................33

6.  Using the 8-Axis Controller...............................33

7.  Appendix: WIRO Prime, an Example..........................35

A.  Data Sheets

1.  IMS #MH2-2222-S Stepping Motor......................3 pp

2.  Berg Drive Chains...................................1 p

3.  Prontor Type 176 Shutter............................2 pp

4.  Guardian #28P-I-24D Solenoid........................1 p

5.  Gordos #832A-1 Relay................................1 p

6.  Micronas #HAL509 Hall Effect Switch................16 pp

B.  Schematics

1.  WIRO Prime Axis Control.............................2 pp

2.  WIRO Prime Relay Board..............................1 p

C.  WIRO Prime Connectors and Circuit Board Layout........1 p

D.  Wiring Procedure: WIRO Prime Ring

Stepper Motor Cable...................................4 pp

E.  Software

1.  WIRO Prime Visual Basic Form........................1 p

2.  WIRO Prime Visual Basic Form As Text................9 pp

3.  WIRO Prime Visual Basic Source Code.................9 pp

WIRO 8-AXIS CONTROLLER

USER'S MANUAL

1.  Purpose of the 8-Axis Controller

The 8-Axis Controller can be incorporated into most stepper               motor systems at WIRO.  The Controller's interface connectors are convenient to the telescope's prime focus and backplate as well as the optics bench. As many as eight steppers up to 3 A peak (4 A with modification to the controller) can be driven by any computer on the network.

1.A.  Why Have an 8-Axis Controller?

WIRO has become involved in the fabrication of instrumentation with adjustments or options that require automation for the sake of efficiency.  Rather than implement separate automation systems for each instrument, a decision has been made to adopt a single controller for most instruments.  These instruments will retain their own actuators and sensors; however, these actuators and sensors will connect to and be driven by the single Controller.

1.B.  What Can this 8-Axis Controller Do?

The Controller is not a complete system; it can do nothing by itself.  Up to eight stepper motors may be connected to the Controller.  Typically, only one motor may move at one time.  Five motors on the backplate, and either five motors on the optics bench, or (eff. Aug04) 3 motors at prime focus may be connected to the controller (subject to the eight motors maximum).

Sensors may be connected to the Controller from both the backplate and the optics bench as follows: two "limits" and one "home" committed to each of four motors (twelve wires total) plus twelve uncommitted "triggers".  Or, instead of connecting from the optics bench, four "triggers" (eff. Aug04) may be connected from prime focus.  These sensors must switch the corresponding 6K8 input between two electrical states, +24 V and ground.  They may be electro-mechanical or solid-state switches.

Four outputs may be driven by the Controller at both the backplate and the optics bench.  Or, instead of connecting to the optics bench, one output (eff. Aug04) may be driven by the Controller at prime focus.  These are open collector outputs which can either provide +24 VDC through 33k ohms or sink 300 mA (max.) to ground.  These outputs may be suitable for energizing relay coils or powering LED's.

The Controller provides power for the sensors and outputs electrically connected to it.  The Controller can also provide power for the general use of the astronomical instrument itself.  +24 VDC (up to 2.4 A and through about 1 ohm) is available at both the backplate and the optics bench; or, instead of connecting to the optics bench, +24 VDC is available at prime focus (through 1.2 ohms).  Also, +5 VDC (up to 3 A) is available at both the backplate (through 1.1 ohms) and the optics bench (through 1.2 ohms); or, instead of connecting to the optics bench, +5 VDC is available at prime focus (through 1.4 ohm).

2.  Scope the User's Manual

The User's Manual describes how to use the 8-Axis Controller (See section 6.).  It also describes the 8-Axis Controller in sufficient detail for an evaluation of its suitability for new projects.  The User's Manual can also be used to evaluate the suitability of modifications to working systems which already make use of the Controller.  Other references, which are listed in section 3., should be consulted before new projects or modifications to existing systems are designed.

The Maintenance Manual for the 8-Axis Controller assumes knowledge of information contained in the User's Manual.  It contains detailed assembly instructions, specification sheets and requisition forms which may be helpful for maintaining and modifying the 8-Axis Controller.

3.  Other Reference Material

For questions which are particular to 8-Axis Controller, as opposed to the IM483 or the 6K8, please consult the Controller's Maintenance Manual and the Observatory Engineer.

For questions concerning the IM483 or the 6K8, please consult the following manufacturer's documentation.  Copies of these should reside in the Control Room at WIRO, in the WIRO Labs, downtown, and the Observatory Engineer's office.

Intelligent Motion Systems (IMS) #IM483 High Performance Microstepping Driver Operating Instructions, Revision 08.02.2000

Parker Automation 6K Series Hardware Installation Guide, Effective: 9Nov98

Parker Automation 6K Series Command Reference, Effective: 5Nov98

Parker Automation 6K Series Programmer's Guide, Effective: 7Nov98

Parker Automation Motion Planner CD, Version 4.1.6

4.  Basic Description

4.A.  Layout and Cabling

This narration is illustrated by the following three diagrams:

4.A.1., 4.A.2. and 4.A.3.

This 8-Axis Controller makes use of one "8-Axis Servo/Stepper Controller" model 6K8 manufactured by Compumotor Division of Parker Hannifin Corporation.  The 6K8 receives commands from a computer via ethernet.  Effective August, 2004, the computer is a Microsoft Windows machine, "WIRO Prime"; and, the control programs are written in Microsoft Visual Basic.  In response to instructions from the computer, the 6K8 can produce either a +/- 10 V signal for servo motors or both step and direction signals for stepper motors; however, the WIRO 8-Axis Controller is wired for stepper motors only; so, servo commands should be avoided (see 5.B.).

The 6K8 has eight Drive connectors which provide eight sets of step and direction signals.  However, these signals cannot drive stepper motors directly.  Between the 6K8 and each stepper motor is a "High Performance Microstepping Driver" model IM483 manufactured by Intelligent Motion Systems, Inc. (IMS).  Thus, the 8-Axis Controller contains eight IM483's.

[The 6K8 also has eight Encoder connectors; but, these are unused in this application.  So, encoder commands should be avoided.]

The 8-Axis Controller box is made of steel which is painted gray and measures 24" by 16" by 8" high.  The 6K8 and the IM483's are mounted within this box along with four power supplies.  The 6K8 front panel LED's are visible through a window in the lid of the box.  Since the operating temperature range of the 6K8 is zero to fifty degrees Celsius, the box is kept in the WIRO Control Room.

Two 48-pin circular connectors are mounted on the box.  These connectors are labeled "Channel 1" and "Channel 2".  They are keyed differently.  Channel 1 carries the four motor drives, twenty-four inputs, and four outputs which are provided by the 6K8 for axes 1 through 4.  Similarly, Channel 2 carries the other four motor drives (axes 5 through 8), their twenty-four inputs, and their four outputs.  A three position switch within the Controller box allows motor 4 to be used on Channel 2 or motor 5 to be used on Channel 1 (Please refer to 4.C.5.).

One permanently routed and dedicated cable runs from the Control room to the telescope backplate (Please refer to 4.C.2.).  A second permanently routed and dedicated cable runs from the Control room to the optics bench which is located inside the Spectrograph room in the Dome (Please refer to 4.C.3.).  Within the Control room, the cable to the backplate can only connect to Channel 2 and is labeled "Channel 2 to Backplate".  The cable to the optics bench can only connect to Channel 1 and is labeled "Channel 1 to Optics Bench".  Both of these cables are used with WIRO's spectrograph, "WIRO Spec", while the telescope is in the cassegrain-focus configuration.

When the telescope is in the prime-focus configuration, alternate cabling is used to bring much of Channel 1's capability to a prime-focus instrument, such as "WIRO Prime".  A permanently routed cable connects the Control room to the top end of the east strut.  (It is used to control the Chopping Secondary Mirror when the telescope is in the cassegrain-focus configuration.  Formerly, it had been unused in the prime-focus configuration.)  Within the Control room this cable terminates at a 100-pin Elco connector which is located on the wall below the breaker panel.  A special-purpose, 68-inch-long adaptor cable connects Channel 1 to this Elco connector.  It is labeled "Channel 1 to Top End".  Thus, the following subset of all the axes-one-through-four signals are brought from the Controller box to the top end of the east strut: motors one through three, trigger inputs two through five, and output four. Please refer to 4.C.4.  [Seven and one half wire pairs remain uncommitted in the Elco connector; however, because of wire gauge limitations, adding motor four would consume six of these pairs.  Refer to 5.A.2.b.1. and 5.A.2.b.5.]

The following drawings illustrate the layout and cabling of the 8-Axis Controller.  They are included next within the User's Manual.

4.A.1.  WIRO 8-Axis Controller Box Layout......................9

4.A.2.  WIRO 8-Axis Controller Cassegrain Focus Cabling.......10

4.A.3.  WIRO 8-Axis Controller Prime Focus Cabling............11

4.B.  Functions

The following documents describe the used functional capability of the IM483 and 6K8 in greater detail.  They are included next within the User's Manual.  Since the IM483 and 6K8 are inside the Controller box, the connectors referred to by these documents are only indirectly accessible to the user.  Please refer to section 4.C. for a description of the 48-pin circular connectors or 100-pin Elco connectors which make these signals available to the user.

4.B.1.  Compumotor 6K8 "Limits/Home" Schematic................13

4.B.2.  Compumotor 6K8 "Triggers/Outputs" Schematic...........14

4.B.3.  Compumotor 6K8 Sample Input Connections...............15

4.B.4.  Compumotor 6K8 Sample Output Connections..............16

4.B.5.  Compumotor 6K Series Commands.........................17

Some commands should be used with care or avoided entirely since certain front panel connections on the 6K8 are left unwired inside the 8-Axis Controller.  Please refer to 5.B.

4.B.6.  IMS Model IM483 Microstepping Driver Quick

Refer Card............................................18

Please see section 4.C.1. for IM483 connections, switch settings and resistor values as implemented in the WIRO 8-Axis Controller.

4.C.  Electrical Configuration

The following documents describe the electrical configuration of the 8-Axis Controller in greater detail.  They are included next, within the User's Manual.

4.C.1.  Typical Motor Drive and Power Supply Configuration....21

4.C.2.  Channel 2 With Telescope Backplate Connector..........22

4.C.3.  Channel 1 With Dome Optics Bench Connector............23

4.C.4.  Telescope Top-End Connector...........................24

4.C.5.  Channel Selector Switch for Motors 4 & 5..............25

The "Channel 1" and "Channel 2" circular connectors which are mounted on the 8-Axis Controller box are wired identically to 4.C.3. and 4.C.2. respectively.

5.  Interfacing the 8-Axis Controller to an Instrument

5.A.  Hardware Issues

5.A.1.  Interface Specifications

5.A.1.a.  Stepper Motor Selection

Stepper motors must conform to motor voltage, peak current, holding current, microstep size, inductance and wiring constraints imposed by the IM483 as implemented within the 8-Axis Controller.  Please refer to sections 4.B.6. and 4.C.1.  Effective August, 2004, for all 8 axes, motor voltage is 24 V, peak current is 3 A (but please see 5.A.2.a. paragraphs 1 and 2), holding current is 0.2 A and microstep size is 2 microsteps per step.

5.A.1.b.  Sensor Selection for "Home", "Limit" or

"Trigger" Inputs

Please refer to sections 4.B.1., 4.B.2., and 4.B.3.  Electrically, the three types of inputs are identical; however, they are distinguished in software.  Please refer to sections 4.B.5. and 5.A.2.a.

5.A.1.c.  "Outputs" Capability

Please refer to sections 4.B.2., 4.B.4. and 5.A.2.a.  When connecting an inductive load to a 6K8 output, it is important to place a flyback diode across the inductance.  See 7.B.2.

5.A.1.d.  Power Capability

Please refer to paragraph 4 of section 1.B. and section 5.A.2.a.  When switching an inductive load across a power supply, it is important to place a flyback diode across the inductance.  See 7.B.2.

5.A.2.  Applications

5.A.2.a.  Mounting Hardware on the Instrument

Stepper Motors designed for peak currents lower than 3.0 A will not necessarily be damaged when used with the 3.0 A peak drive provided by the 8-Axis Controller.  Damage to most motors will not occur if the motor's internal temperature is kept below 130 Celsius.  Assuming a 30 degree gradient, the motor's outside case must be kept below 100 Celsius.  As a guide, metal hot enough to touch without burning, but too hot to remain in contact with, is at about 65-70 Celsius.  Since the holding current is 0.2 A, reduced duty cycles should allow the safe operation of many stepping motors with peak currents lower than 3.0 A.

Stepper motors with peak currents higher than 3.0 A may work well with the 8-Axis Controller depending upon the application.  If the application requires more than 3.0 A for dependable operation then the "current adjust resistor" on the corresponding IM483 should be changed.  Please consult the maintenance manual and refer to 4.B.6. before attempting this modification.  The IM483 cannot provide more than 4.0 A peak.

Stepper motor control of an axis of motion will usually require additional sensors or actuators.  Snap-acting switches, Hall-effect sensors, photoelectric sensors, or fiber-optic sensors may be used to detect particular positions upon the axis.  A solenoid may be used to lock the position of the axis.  The WIRO 8-Axis Controller provides +24 VDC which can be used to power the sensors and actuators.  Please refer to 4.C.2., 4.C.3. and 4.C.4.

Position detection can be accomplished through the dedicated use of a single "Trigger" input; or, if multiple positions must be detected, then a binary encoding scheme can detect up to (2^X - 1) positions with just (X) inputs. Commands issued over the network to the 6K8 can read the state of one or many "Trigger" inputs.  Please refer to sections 4.B.2., 4.B.3. and 7.B.1.

Special positions for initializing motion or for limiting the range of motion should make use of the "Home" and "Limit" inputs.  These inputs are used by many useful 6K8 algorithms which anticipate typical problems in homing or limiting motion.  These algorithms are invoked with simple commands to the 6K8 via the network.  Please refer to sections 4.B.1., 4.B.3. and 4.B.5.

[The Compumotor 6K8 has separate connectors for each axis which could be used to implement position feedback by incremental encoder; however, these eight connectors are unused in the WIRO 8-Axis Controller.  Additionally, each of the axes could be configured to provide +/- 10 volt drive signals for servo motor control of motion; however, only the step and direction signals for stepping motors are used in the WIRO 8-Axis Controller.  Thus, encoders and servo motors cannot currently be used with the WIRO 8-Axis Controller; and, all 6K8 commands for encoders and servos should not be invoked.  Please refer to 5.B.]

Separate actuators may be necessary for assisting with stepper motor control of an axis.  The 6K8 "Outputs" can energize such actuators directly, by sinking not more than 300 mA, or indirectly by means of a relay.  Power can be provided by the Controller's 24 VDC supply for external functions.  Commands are issued over the network in order to change the state of the "Outputs".  Please refer to sections 4.B.2., 4.B.4., 4.C.1. and 7.B.2.

"Triggers" and "Outputs" from the 6K8 may also be used to implement functions which are unrelated to the control of motion along an axis which is propelled by a stepper motor.  For example, WIRO Spec uses four separate outputs to turn on argon, xenon, neon, and incandescent calibration sources.

+24 VDC from the WIRO 8-Axis Controller can be used to power actuators which are controlled by a 6K8 output via a relay.  For example, the filter wheel solenoid on WIRO Prime draws 0.5 A at 24 V while the 6K8 output which controls the relay need only provide 25 mA.  Also, actuator power need not be controlled by a 6K8 output.  For example, WIRO Prime's 24 V shutter current of 0.9 A is switched on and off by an output from the Leach camera controller.  Please refer to 7.B.1., 7.B.2., 7.A.5. and 7.A.3.

5.A.2.b.  Cabling the Controller to the Instrument

5.A.2.b.1.  Wire Gauge and Length Restrictions

Stepper motor wiring should not exceed a length of 100 feet from the 8-Axis Controller; and, it should be 18 gauge.  Please refer to 4.C.2., 4.C.3. and 4.C.4.

Inputs to and outputs from the 6K8 are not restricted as to length or wire gauge.

Delivery of +24 V power and +5 V power should account for the IR voltage drop which will occur when the required current (I) flows through wiring which has an inherent resistance (R).  Round trip resistances to the backplate, optics bench, and top end are given in paragraph four of section 1.B.  Additional cabling from these positions to the current consuming device will introduce additional resistance.  For example, 22 gauge has .0165 ohms/ft, 19 gauge has .0082 ohms/ft and 16 gauge has .0041 ohms/ft.  18 gauge is .0065 ohms/ft.  Select a wire gauge which yields an acceptable loss of voltage.

5.A.2.b.2.  Twisted Pair Wiring

Minimize radiation from rapidly changing currents by routing an outbound current with its return current if possible.  Use shielded, twisted-pair wiring.  Please refer to 5.A.2.b.4., 4.C.4. and 5.A.2.b.5.

5.A.2.b.3.  Shielding

Much effort has been expended in order to properly shield 8-Axis Controller cabling.  In the case of stepper motor wiring and +24 V power wiring, the Controller chassis was initially wired to the shielding via the box's circular connector bodies and screw lugs (as opposed to using one of the 48 pins).  Dedicated to this purpose, these shields were left unconnected at their other ends (as is standard practice) at the backplate, optics bench and top end.  However, random motor motion was encountered in the first tests of these cables.  At the suggestion of IMS technical support, the lugs were unscrewed from the connector bodies so that the shielding was left to float in an electrical sense.  This simple step eliminated the random motor motion; so, the lugs remain unscrewed and are currently insulated with electrical tape.

Unshielded motor cabling is not advised.  Prototype cabling to the optics bench was initially unshielded; when, the camera's readout picked up tremendous amounts of noise.  This noise was caused by simply turning "on" the 8-Axis Controller; no motion was commanded at any time during the exposure or readout.  In this case, noise on the camera readout was eliminated by (1) insulating the camera from the optics bench, (2) wrapping the four motor cables collectively in a couple layers of aluminum foil and (3) connecting the aluminum foil to the optics bench itself.  This is the final electrical configuration of the motor shields on the short cable which joins the circular connector to the optics bench itself (except that the motor phases are shielded individually).

All 6K8 Trigger, Home, Limit and Output signals are shielded by the 6K8's Isolated Ground.  "Iso Gnd" occupies pin z in the circular connectors or pins T2, T10 and U7 in the Elco connectors.  Iso Gnd should be used as a reference or return path for these 6K8 I/O signals as shown in 4.B.3. and 4.B.4.  Thus, 6K8 I/O signal wiring resembles coaxial wiring except that the shield surrounds a twisted pair.  Wire pairing of these 6K8 I/O signals within the Iso Gnd shields is somewhat arbitrary.  Please refer to 4.C.4. and the following two diagrams, included next within the User's Manual:

5.A.2.b.4.  Wiring Diagram for Circular Connectors............31

5.A.2.b.5.  Wiring Diagram for Elco Connectors................32

5.B.  Software Issues

Where software is concerned, use of the WIRO 8-Axis Controller requires (1) the understanding of Compumotor 6K8 commands and (2) the creation of algorithms which make use of the 6K8 commands.  Please refer to section 3.  These algorithms can be programmed into any computer that can use the network to address the 6K8 (refer to 4.A.1.).  Microsoft support for Compumotor is more widely disseminated than others; and, currently WIRO makes use of the Windows machine, "WIRO Prime" to command the 8-Axis Controller via Visual Basic.  Software for the instrument WIRO Prime is in section 7.E. as an example.

The WIRO 8-Axis Controller does not use connections on the 6K8 front panel which are intended for incremental encoders or servo motors.  Therefore, all commands which reside under the functional groupings "Encoder" and "Servo" on 4.B.5. should be used with special caution.  Also, the "6K Series Command Reference" description of the AXSDEF command lists the following commands as "servo only commands": ANIFB, KDRIVE, PER, SFB, SGP, SGI, SGV, SGVF, SGAF, SGILIM, SOFFS, SMPER, SGSET, SGENB, STRGTD, STRGTE, STRGTT, STRGTV, TFB, TGAIN, TPER, TSGSET and TSTLT.

6K8 commands related to "shutdown" may be essential for obtaining zero motor motion.  In particular, motors which have received no command to move can still vibrate and even drift in position unless the "shutdown" command has been issued.  "Shutdown" makes use of the connection between pin 12 on the 6K8 Drive connectors, "Shutdown-", and pin 5 on the IM483 P1 connectors, "Enable/Disable Input".  Please refer to 4.C.1.

Keep in mind that different models of stepping motors may have different step sizes.  Also, all eight of the IM483's are currently NOT set to 2 microsteps per step as shown in the schematic.  Please refer to 4.C.1.

6.  Using the 8-Axis Controller

Verify that all the proper cables are plugged into the Controller box.  A power cable and an Ethernet cable should be plugged into the rear panel.  For prime-focus instruments the cable labeled "Channel 1 to Top End" should be plugged into the Channel 1 connector on the front panel.  For cassegrain-focus instruments the cable labeled "Channel 2 to Backplate" should be plugged into the Channel 2 connector on the front panel.  For cassegrain-focus spectrographs both the "Channel 2 to Backplate" and the "Channel 1 to Optics Bench" cables should be plugged into the front panel.  Make sure that the instrument is properly plugged into the circular connectors at the backplate and/or optics bench or the Elco connector at the top end of the east strut.

Verify that the three position switch is set properly.  The setting can be checked by lifting the cover and examining the switch itself (front panel LED's are not yet implemented).  The center setting sends motor 4 to Channel 1 and motor 5 to Channel 2.  The left setting sends both motors 4 and 5 to Channel 1.  The right setting sends both motors 4 and 5 to Channel 2.

Turn on the power switch on the rear panel of the 8-Axis Controller box.  Through the window on the lid of the 8-Axis Controller box the eight red axis lights and the green power light should be visible on the front panel of the 6K8.  If the power light is red, flip the power switch off and on again.  If the power light remains red, call the observatory engineer.  Wait several seconds while one of the LED's on the 6K8's Ethernet connector blinks on-off-on once.

Log onto the computer named "WIRO Prime" and activate the shortcut to the desired instrument controller software.  A graphical user interface (GUI) should appear on the monitor.  Click on the "connect" button with the mouse.  After several seconds a red disk on the GUI should turn to green indicating that the network connection between the computer and the 6K8 has been established.

Within the GUI click on a button labeled something like "initialize axes".  Look through the window in the lid of the box and verify that certain of the eight red axis lights have been extinguished.  Each extinguished lamp corresponds to one of the instrument motors.  Now the instrument motors, outputs, and sensors can be controlled by clicking on GUI buttons.  If a power consuming device other than a stepper motor (eg: solenoid or shutter) is activated then the front panel current meter will display the 24 V current which is being drawn by the device.

To turn off the 8-Axis Controller first click on the GUI "disconnect" button.  The green disk on the GUI should turn red.

Now the power switch on the rear panel of the box can be turned off and the GUI window can be closed.

7.  Appendix:  WIRO Prime, an Example

The instrument WIRO Prime is described on a web page which can be accessed through our department's website, "http://physics.uwyo.edu", and following links labeled "Observatories", "Wyoming Infrared Observatory", and "WIRO-Prime".  Effective September, 2004, this web page can be accessed directly at "http://physics.uwyo.edu/~mpierce/WIRO/Prime.html".

The following information pertains to the way in which WIRO Prime utilizes the WIRO 8-Axis Controller; and, it can be used as an example for the use of the Controller by an instrument:

pages

7.A.  Data Sheets

7.A.1.  IMS #MH2-2222-S Stepping Motor.........................3

7.A.2.  Berg Drive Chains......................................1

7.A.3.  Prontor Type 176 Shutter...............................2

7.A.4.  Guardian #28P-I-24D Solenoid...........................1

7.A.5.  Gordos #832A-1 Relay...................................1

7.A.6.  Micronas #HAL509 Hall Effect Switch...................16

7.B.  Schematics

7.B.1.  WIRO Prime Axis Control................................1

7.B.2.  WIRO Prime Relay Board.................................1

7.C.  WIRO Prime Connectors and Circuit Board Layout...........1

7.D.  Wiring Procedure: WIRO Prime Ring Stepper Motor Cable....4

7.E.  Software

7.E.1.  WIRO Prime Visual Basic Form...........................1

7.E.2.  WIRO Prime Visual Basic Form As Text...................9

7.E.3.  WIRO Prime Visual Basic Global Module..................2

7.E.4.  WIRO Prime Visual Basic Source Code...................15

'============================================================================

'

' FILE:       GlobalModule.BAS

'

' PURPOSE:    Global declarations for VB example using Com6srvr.exe

'

' COPYRIGHT:  Copyright(c) 1992-2000, All Rights Reserved

'             Parker Hannifin Corporation

'             Compumotor Corporation

'             5500 Business Park Drive

'             Rohnert Park, California 94928

'             Applications Engineering:  (800)358-9070

'

' DISCLAIMER: THIS SOFTWARE IS PROVIDED FREE OF CHARGE AND WITHOUT

'             WARRANTY OF ANY KIND, EITHER EXPRESSED OR IMPLIED.  IN NO

'             EVENT WILL PARKER HANNIFIN CORPORATION BE LIABLE FOR ANY

'             DAMAGES, INCLUDING ANY LOST PROFITS, LOST SAVINGS, OR OTHER

'             INCIDENTAL OR CONSEQUENTIAL DAMAGES ARISING OUT OF THE USE

'             OR INABILITY TO USE THIS SOFTWARE.

'

' NOTE:       Additional information regarding the Com6srvr.exe driver is

'             avaliable in the readme.txt file in the Com6srvr directory.

'

'============================================================================

Option Explicit

'------------------------------------------------------------------------------------------------------------------------------------------

' Faststatus Information Data Structure

'------------------------------------------------------------------------------------------------------------------------------------------

Type FastStatusInfo

UpdateID As Integer               ' Reserved for internal use

Counter As Integer                ' time frame counter (2ms per count)

MotorPos(1 To 8) As Long          ' commanded position (counts)

EncoderPos(1 To 8) As Long        ' actual position (counts)

MotorVel(1 To 8) As Long          ' commanded velocity (counts/sec)

AxisStatus(1 To 8) As Long        ' axis status (TAS)

SysStatus As Long                 ' system status (TSS)

ErrorStatus As Long               ' user status (TER)

UserStatus As Long                ' user status (TUS)

Timer As Long                     ' timer value (TIM - milliseconds)

Limits As Long                    ' limit status (TLIM)

ProgIn(0 To 5) As Long            ' programmable input status (TIN)

ProgOut(0 To 3) As Long           ' programmable output status (TOUT)

Triggers As Long                  ' trigger interrupt status (TTRIG)

Analog(1 To 2) As Integer         ' lo-res analog input voltage (TANV)

VarB(1 To 10) As Long             'VARB1 - VARB10

VarI(1 To 10) As Long             'VARI1 - VARI10

Reserved As Long                  'Reserved for internal use

CmdCount As Long                  'Command Count (from communications port)

End Type

'------------------------------------------------------------------------------------------------------------------------------------------

' Win API Function declarations

'------------------------------------------------------------------------------------------------------------------------------------------

Public Declare Sub CopyMemory Lib "Kernel32" Alias "RtlMoveMemory" (destination As Any, source As Any, ByVal numbytes As Long)

'------------------------------------------------------------------------------------------------------------------------------------------

' Global variables

'------------------------------------------------------------------------------------------------------------------------------------------

Global c6k As Object            ' comm server object (RS232 or Ethernet)

Global fsinfo As FastStatusInfo 'fast status information

Global connected As Boolean     ' flag to indicate connection state - could also use the state of timer1.enabled for this

Global FltList(12) As String

Global FltPos As Integer

Global FocusPos As Integer

Global CorrPos As Integer

Global CorrMin As Integer

Global CorrMax As Integer

Global FocusMin As Integer

Global FocusMax As Integer

'updated 08/01/03

Option Explicit

Dim i As Integer

Dim FiltPos As Single       ' FiltPos is the integer position of current filter and is globally defined

Dim NewPos As Single        ' NewPos is the integer position of where we go

Dim NewFilt As String

'Dim FltPos As Single

Dim FiltFree As Integer

Dim FiltStps As Single

Private Sub Lites()

If CStr(fsinfo.ProgIn(0)) = 0 Then FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbRed: FiltMag4.BackColor = vbRed

If CStr(fsinfo.ProgIn(0)) = 2 Then FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbRed: FiltMag4.BackColor = vbGreen

If CStr(fsinfo.ProgIn(0)) = 4 Then FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbGreen: FiltMag4.BackColor = vbRed

If CStr(fsinfo.ProgIn(0)) = 6 Then FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbGreen: FiltMag4.BackColor = vbGreen

If CStr(fsinfo.ProgIn(0)) = 8 Then FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbGreen: FiltMag3.BackColor = vbRed: FiltMag4.BackColor = vbRed

If CStr(fsinfo.ProgIn(0)) = 10 Then FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbGreen: FiltMag3.BackColor = vbRed: FiltMag4.BackColor = vbGreen

If CStr(fsinfo.ProgIn(0)) = 12 Then FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbGreen: FiltMag3.BackColor = vbGreen: FiltMag4.BackColor = vbRed

If CStr(fsinfo.ProgIn(0)) = 14 Then FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbGreen: FiltMag3.BackColor = vbGreen: FiltMag4.BackColor = vbGreen

If CStr(fsinfo.ProgIn(0)) = 16 Then FiltMag1.BackColor = vbGreen: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbRed: FiltMag4.BackColor = vbRed

If CStr(fsinfo.ProgIn(0)) = 18 Then FiltMag1.BackColor = vbGreen: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbRed: FiltMag4.BackColor = vbGreen

If CStr(fsinfo.ProgIn(0)) = 20 Then FiltMag1.BackColor = vbGreen: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbGreen: FiltMag4.BackColor = vbRed

If CStr(fsinfo.ProgIn(0)) = 22 Then FiltMag1.BackColor = vbGreen: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbGreen: FiltMag4.BackColor = vbGreen

If CStr(fsinfo.ProgIn(0)) = 24 Then FiltMag1.BackColor = vbGreen: FiltMag2.BackColor = vbGreen: FiltMag3.BackColor = vbRed: FiltMag4.BackColor = vbRed

End Sub

Private Sub cmdConnect_Click()

Dim iConnectStatus As Integer

Dim cmdrtn As Long

Dim sIpAddr As String

sIpAddr = txtIpAddr.Text  ' Get the IP address ' Warning! -- No validation of Ip address format done here

On Error GoTo err_connect_handler

If Not connected Then  ' If a connection has not been established make one

Set c6k = CreateObject("COM6SRVR.NET")  ' Create the Com6srvr

iConnectStatus = c6k.Connect(sIpAddr)  ' Connect to the 6k

' Connect method returns positive number if successful

' if negative number returned connect attempt was unsuccessful

If iConnectStatus > 0 Then

connected = True                  'set connected flag to true

c6k.FSUpdateRate = 100            'set fast status update rate to 100ms

c6k.FSEnabled = True              'enable fast status polling

shpTrafficLight.FillColor = vbGreen

Else

connected = False           'set connected flag to false

Set c6k = Nothing           ' Destroy Com6srvr object

MsgBox "Connection attempt failed..." & CStr(iConnectStatus), 0, "Status"

End If

End If

Wait (3)

Call FiltRead_Click

Exit Sub  ' exit sub now don't go into error handler

err_connect_handler:

Call Disconnect

End Sub

Private Sub cmdDisconnect_Click()  ' Call the disconnect procedure to disconnect

Open "c:\FiltPos.txt" For Output As #1

Write #1, FiltPos

Close #1

Call SolEng_Click

Call Disconnect

End Sub

Private Sub Command1_Click()

Dim Temp() As Byte

If connected Then

Temp = c6k.FastStatus

Call CopyMemory(fsinfo, Temp(0), 280)

Text3.Text = fsinfo.ProgIn(0)

Text2.Text = dec2bin(fsinfo.ProgIn(0))

Picture1.Print fsinfo.ProgIn(0)

End If

End Sub

Public Function dec2bin(mynum As Variant) As String

Dim Loopcounter As Integer

If mynum >= 2 ^ 31 Then

dec2bin = "Too Big"

Exit Function

End If

Do

If (mynum And 2 ^ Loopcounter) = 2 ^ Loopcounter Then

dec2bin = "1" & dec2bin

Else

dec2bin = "0" & dec2bin

End If

Loopcounter = Loopcounter + 1

Loop Until 2 ^ Loopcounter > mynum

End Function

Private Sub FiltJogM_Click()

Call SolDis_Click

'Dim FiltStps As Single

c6k.Write "D,,-5:"

c6k.Write "GO001:"

FiltStps = FiltStps - 5

FiltSteps.Text = FiltStps

Wait (2)

Call SolEng_Click

Call FiltRead_Click

End Sub

Private Sub FiltJogP_Click()

Call SolDis_Click

'Dim FiltStps As Single

c6k.Write "D,,5:"

c6k.Write "GO001:"

FiltStps = FiltStps + 5

FiltSteps.Text = FiltStps

Wait (2)

Call SolEng_Click

Call FiltRead_Click

End Sub

Private Sub Initialize_Click()

Dim sResponse As String

If connected Then

c6k.Write "TREV:"  ' Send the 6K TREV command ' Note -- Colon character is delimiter

c6k.Write "DRIVE0000:"

c6k.Write "AXSDEF0000:"

c6k.Write "LH0,0,0,0:"

c6k.Write "DRFLVL0:"

c6k.Write "DRFEN1:"

c6k.Write "DRIVE1110:"

c6k.Write "MC0,0,0:"

c6k.Write "SCALE1:"

c6k.Write "SCLD1,1,1:"

c6k.Write "SCLV1,1,1:"

c6k.Write "SCLA1,1,1:"

c6k.Write "A50,50,100:"

c6k.Write "V100,100,300:"

End If

End Sub

Private Sub CorrMinus_Click()

If CorrPos - 50 >= CorrMin Then

c6k.Write "D,-50:"

c6k.Write "GO01:"

CorrPos = CorrPos - 50

Corrector.Text = CStr(CorrPos)

Else

MsgBox "Minimum Corrector Position Exceeded!"

End If

End Sub

Private Sub CorrPlus_Click()

If CorrPos + 50 <= CorrMax Then

c6k.Write "D,50:"

c6k.Write "GO01:"

CorrPos = CorrPos + 50

Corrector.Text = CStr(CorrPos)

Else

MsgBox "Maximum Corrector Position Exceeded!"

End If

End Sub

Public Sub FilterGo_Click()

Call SolDis_Click

If FiltFree = 1 Then

Dim i As Integer

Dim Change As Single    ' Change is the number of filter positions we need to move

Dim Command As String   ' Command is the command string to be sent to the c6k

Dim StrSteps As String

Dim ScaleFact As Integer

Dim Steps As Integer

'    ScaleFact = 334.3            ' Set the scale factor of # of steps/ filter pos.

ScaleFact = 332.5            ' Set the scale factor of # of steps/ filter pos.

NewFilt = Filter.Text    ' Read the filter name from the filter combo box

For i = 1 To 12

If NewFilt = FltList(i) Then

NewPos = i

End If

Next i

Change = NewPos - FiltPos       ' Compute the offset in filter units

Steps = ScaleFact * Change      ' Compute the number of steps using an optional scale factor

StrSteps = CStr(Steps)          ' Convert the number of steps to a string

Command = "D,," + StrSteps + ":"

c6k.Write Command               ' Issue the command to the c6k

c6k.Write "GO001:"              ' Issue the command to the c6k

FiltPos = NewPos                ' Update the current filter position

Open "c:\FiltPos.txt" For Output As #1

Write #1, FiltPos

Close #1

Else: MsgBox "Unlock Filter Wheel!"  'remove?

End If

Wait (3 * Abs(Change))

Call SolEng_Click

Wait (2)

Call FiltRead_Click

End Sub

Private Sub FiltRead_Click()

Dim Temp() As Byte            'create temporary byte array

If connected Then             'Copy the c6k>FastStatus packet into a byte array

Temp = c6k.FastStatus                   'get fast status information

'Use CopyMemory to map the byte array into the fsinfo structure

Call CopyMemory(fsinfo, Temp(0), 280)       'copy byte array into structure

FiltNumb.Text = ""

Text1.Text = CStr(fsinfo.ProgIn(0))

FiltNumb.Text = FiltNumb.Text & CStr(fsinfo.ProgIn(0))

If CStr(fsinfo.ProgIn(0)) = 0 Then FiltPos = 0

If CStr(fsinfo.ProgIn(0)) = 2 Then FiltPos = 1

If CStr(fsinfo.ProgIn(0)) = 4 Then FiltPos = 2

If CStr(fsinfo.ProgIn(0)) = 6 Then FiltPos = 3

If CStr(fsinfo.ProgIn(0)) = 8 Then FiltPos = 4

If CStr(fsinfo.ProgIn(0)) = 10 Then FiltPos = 5

If CStr(fsinfo.ProgIn(0)) = 12 Then FiltPos = 6

If CStr(fsinfo.ProgIn(0)) = 14 Then FiltPos = 7

If CStr(fsinfo.ProgIn(0)) = 16 Then FiltPos = 8

If CStr(fsinfo.ProgIn(0)) = 18 Then FiltPos = 9

If CStr(fsinfo.ProgIn(0)) = 20 Then FiltPos = 10

If CStr(fsinfo.ProgIn(0)) = 22 Then FiltPos = 11

If CStr(fsinfo.ProgIn(0)) = 24 Then FiltPos = 12

'If FltPos = FiltPos Then

If FiltPos > 0 Then

NewFilt = FltList(FiltPos)

FiltNumb.Text = FiltPos & "  " & NewFilt

Filter.Text = NewFilt

Call Lites

Else: Call Lites2

'MsgBox "Initialize Filter Wheel!"

End If

End If

Open "c:\FiltPos.txt" For Output As #1

Write #1, FiltPos

Close #1

End Sub

Private Sub FocusMinus_Click()

If FocusPos - 50 >= FocusMin Then

c6k.Write "D-100:"

c6k.Write "GO1:"

FocusPos = FocusPos - 50

Focus.Text = CStr(FocusPos)

Else

MsgBox "Minimum Focus Position Exceeded!"

End If

End Sub

Private Sub FocusPlus_Click()

If FocusPos + 50 <= FocusMax Then

c6k.Write "D100:"

c6k.Write "GO1:"

FocusPos = FocusPos + 50

Focus.Text = CStr(FocusPos)

Else

MsgBox "Maximum Focus Position Exceeded!"

End If

End Sub

Private Sub Form_Load()

connected = False       ' When the form loads (Application starts) we're not connected to 6K

'FltPos = 1             ' Initialize the filter at position #1 at startup

Open "C:\FiltPos.txt" For Input As #1

Input #1, FiltPos

Close #1

Open "C:\Filters.txt" For Input As #1   ' Read the filter list file and store in the FltList Array

Dim j As Integer

For j = 1 To 12

Input #1, FltList(j)

Filter.AddItem FltList(j)

Next j

Close #1

Filter.ListIndex = 0

FocusPos = 0                 ' Initialize the focus and corrector positions at 0

Focus.Text = CStr(FocusPos)

CorrPos = 0                  ' Initialize the focus and corrector positions at 0

Corrector.Text = CStr(CorrPos)

' Seup Corrector and Focus Limits

CorrMin = -10000

CorrMax = 10000

FocusMin = -10000

FocusMax = 10000

FiltFree = 0

Locked.Visible = True

Unlocked.Visible = False

End Sub

Private Sub Form_Unload(Cancel As Integer)

Open "c:\FiltPos.txt" For Output As #1

Write #1, FiltPos

Close #1

Call Disconnect

End Sub

Public Sub Disconnect()

shpTrafficLight.FillColor = vbRed

'Call SolEng_Click?

' If connected stop fast status update by 6k

If connected Then

c6k.FSEnabled = False

End If

' reset the global connected flag

connected = False

' Destroy com6srvr object

Set c6k = Nothing

End Sub

Private Sub InitWhl_Click()

Call FiltRead_Click

Dim Command As String

Dim StrSteps As String

Dim Steps As Single

Dim x As Integer

Dim y As Single

Dim z As Single

y = 5

'y = 37.5

z = 0

Open "c:\FiltPos.txt" For Input As #1

Input #1, FiltPos

Close #1

If 1 <= FiltPos < 12 Then

x = 1

Else: x = -1

End If

If CStr(fsinfo.ProgIn(0)) > 0 Then

Call FiltRead_Click

Else

Call SolDis_Click

If FiltFree = 1 Then

Do

Steps = x * (y + z)         ' Compute the number of steps using an optional scale factor

StrSteps = CStr(Steps)      ' Convert the number of steps to a string

Command = "D,," + StrSteps + ":"

c6k.Write Command           ' Issue the command to the c6k

c6k.Write "GO001:"          ' Issue the command to the c6k

z = (z + y)

x = (-1 * x)

Wait (4)

Call FiltRead_Click

Loop While CStr(fsinfo.ProgIn(0)) = 0

End If

End If

'Else: MsgBox "Unlock FilterWheel"  'remove?

Open "c:\FiltPos.txt" For Output As #1

Write #1, FiltPos

Close #1

Call SolEng_Click

End Sub

Private Sub SolDis_Click()

'Disengage Solenoid

'*Change from 4 to 8 for 6k8 on mountain*??

c6k.Write "OUT.4-1:"

Solenoid.FillColor = vbGreen

Locked.Visible = False

Unlocked.Visible = True

FiltFree = 1

End Sub

Private Sub SolEng_Click()

'Engage Solenoid

'*Change from 4 to 8 for 6k8 on mountain*??

c6k.Write "OUT.4-0:"

Solenoid.FillColor = vbRed

Locked.Visible = True

Unlocked.Visible = False

FiltFree = 0

End Sub

Private Sub Lites2()

FiltMag1.BackColor = vbRed: FiltMag2.BackColor = vbRed: FiltMag3.BackColor = vbRed: FiltMag4.BackColor = vbRed

End Sub

Sub Wait(ByVal s As Integer)

Dim t As Double

t = Now() + s / 86400#

While t > Now()

DoEvents

Wend

End Sub